## Sheet1
Pressure Calibration Certification
|  |  | CUSTOMER | Applied Materials
| PART # |  |  | E28D-31351
| WIP# |  |  | D4615
| DATE |  |  | 2025-09-12
| MODEL |  |  | E28D-31351
| S/N |  |  | 112568239
| RANGE |  |  | 0.1 TORR
| FITTING |  |  | 4fVCR
| CONNECTOR |  |  | 9 PIN
| CAL STD ID |  |  | PVS 100
| MOUNTING ATTITUDE |  |  | Vertical
| TECH ID |  |  | TH
| SERVICE TYPE |  |  | Level 1 Repair
Voltage Set Point |  | STD Gauge Output VDC | DUT Output Voltage |  | Error mVolt | Error % Rdg
10 |  | -0.00404 |  | -0.013153
9 |  | 0.99892 |  | 1.0062
8 |  | 1.999 |  | 2.0033
7 |  | 2.9993 |  | 3.0064
6 |  | 3.9999 |  | 4.0055
5 |  | 4.9996 |  | 5.0037
4 |  | 5.9994 |  | 6.0046
3 |  | 6.9998 |  | 7.006
2 |  | 7.9983 |  | 8.0048
1 |  | 8.9985 |  | 9.0069
0 |  | 9.999 |  | 10.013
|  | Time | Voltage Reading [mv]
|  | 0 Hours | 0 mv
|  | 12 Hour | 1mv PASS (+/- 5 mv)
All Calibrations are performed in a controlled environment by qualified personnel using instrumentation and methods which guarantee that the specifications claimed are reliable. The above data was measured on a primary standard traceable to the National Institute of Standards and Technology and conform to ANSI/NCSL Z540-1 (formerly MILSTD 45662A) & ISO-17025. Contamination, mis-use, abuse or mis-handling voids the warranty in its' entirety. Cal Std Model: PVS6-ID S/N: 000513425