## Sheet1
Pressure Calibration Certification
CUSTOMER
| PART # |  |  | 1350-00856
WIP#
DATE
| MODEL |  |  | E28D-31351
S/N
| RANGE |  |  | 0.1 TORR
| FITTING |  |  | 8fVCR
| CONNECTOR |  |  | 15 PIN
| CAL STD ID |  |  | PVS100
| MOUNTING ATTITUDE |  |  | Vertical
TECH ID
| SERVICE TYPE |  |  | Level 1 Repair
Voltage Set Point |  | STD Gauge Output VDC | DUT Output Voltage |  | Error mVolt | Error % Rdg
10 |  | 10.097 |  | 10.076 | 21 | 0%
9 |  | 9.107 |  | 9.076 | 31 | 0%
8 |  | 8.081 |  | 8.046 | 35 | 0%
7 |  | 7.098 |  | 7.063 | 35 | 0%
6 |  | 6.11 |  | 6.082 | 28 | 0%
5 |  | 5.109 |  | 5.089 | 20 | 0%
4 |  | 4.004 |  | 3.997 | 7 | 0%
3 |  | 3.055 |  | 3.063 | -8 | 0%
2 |  | 2.024 |  | 2.043 | -19 | 0%
1 |  | 1.088 |  | 1.112 | -24 | 0%
0 |  | 0 |  | 0 | 0 | 0%
|  | Time | Voltage Reading [mv]
|  | 0 Hours | 0 mv
|  | 12 Hour | 1mv PASS (+/- 5 mv)
All Calibrations are performed in a controlled environment by qualified personnel using instrumentation and methods which guarantee that the specifications claimed are reliable. The above data was measured on a primary standard traceable to the National Institute of Standards and Technology and conform to ANSI/NCSL Z540-1 (formerly MILSTD 45662A) & ISO-17025. Contamination, mis-use, abuse or mis-handling voids the warranty in its' entirety. Cal Std Model: PVS6-ID S/N: 000513425